Select the correct detector;

In this example, we select

#define HAMA_S11510_1024BIN 11

#define DEETECTOR HAMA_S11510_1024BIN

In xspec_control_rs232, you need to change these parameters for your detector.

if(DETECTOR==HAMA_S8378_256)

{

data_start=S8378_DATA_START;

data_end=S8378_DATA_END;

pixels=S8378_TOTAL_PIXELS;

}

else if(DETECTOR==TCD1304)

{

data_start=T_DATA_START;

data_end=T_DATA_END;

pixels=T_TOTAL_PIXELS;

}

else if(DETECTOR== HAMA_S11510_1024BIN)

{

data_start= S11510BIN_DATA_START;

data_end= S11510BIN_DATA_END;

pixels= S11510BIN_TOTAL_PIXELS;

}

// HAMA S11510BIN 1024 pixel; parameters for S11520BIN

#define S11510BIN_DARK_START 32

#define S11510BIN_DARK_END 33

#define S11510BIN_TOTAL_DARK 2

#define S11510BIN_DATA_START 37//spec says 32

#define S11510BIN_DATA_END 1060

#define S11510BIN_TOTAL_DATA 1024  //3683

#define S11510BIN_TOTAL_PIXELS 1034 //including data dark and some others.

#define INT_S10420_TIME_NORMAL_MIN 9400 //smallest integration time for HAMA_S11510_1024BIN

Read integration time in spectrometer:

comport=XSPEC_rs232_open(devnum);

int temp_int_time=XSPEC_rs232_getIntegrationTime(comport);

XSPEC_rs232_close(comport);

Change integration time in spectrometer:

comport=XSPEC_rs232_open(devnum);

XSPEC_rs232_setIntegrationTimeFlash(comport, int_time);

XSPEC_rs232_close(comport);

Read the calibration data.

SYSTEM_SETTING_COEF cal_coef;

comport=XSPEC_rs232_open(devnum);

XSPEC_rs232_get_SpecConfig0123(comport, (char *)&cal_coef, 2);

XSPEC_rs232_close(comport);

Read a spectrum from spectrometer:

comport=XSPEC_rs232_open(devnum);

bytesread=XSPEC_rs232_measure(comport, data_array, pixels*2+2, 0x1234, temp_int_time);

XSPEC_rs232_close(comport);

From the measured data, you need to convert the pixel number to wavelength. See the print_data() in xspec_control.cpp for example.

You can use XSPEC_rs232_measure_partial() to read part of the spectrum since RS232 is slow.

See below for function parameters.

//open the XSPEC device by comm port number 0,1,2,3,4...

extern "C" INTERFACE_API HANDLE XSPEC_rs232_open(unsigned char devnum);

//call this to close the rs232. 0==success, < 0 on error.

extern "C" INTERFACE_API int XSPEC_rs232_close(HANDLE comport);

// set the integration time in usec for the current_devnum in RAM only, not persistent

// return the number of bytes sent for the command, must be larger than 0.

extern "C" INTERFACE_API int XSPEC_rs232_setIntegrationTime(HANDLE comport,unsigned int time_us);

// set the integration time in usec for the current_devnum, stays in FLASH

// return the number of bytes sent for the command, must be larger than 0.

extern "C" INTERFACE_API int XSPEC_rs232_setIntegrationTimeFlash(HANDLE comport,unsigned int time_us);

// set the integration mode for the current_devnum, stays in FLASH

// mode 0: normal. software

// mode 1: rising external edge

// mode 2: falling external edge

// return the number of bytes sent for the command, must be larger than 0.

extern "C" INTERFACE_API int XSPEC_rs232_setModeFlash(HANDLE comport,unsigned int mode);

// set the firmware averaging for the current_devnum, stays in FLASH

// return the number of bytes sent for the command, must be larger than 0.

extern "C" INTERFACE_API int XSPEC_rs232_setAverageFlash(HANDLE comport,unsigned int average);

// return the integration time in usec for the current_devnum

// -1==failure,

extern "C" INTERFACE_API int XSPEC_rs232_getIntegrationTime(HANDLE comport);

// return the firmware average number.

// -1==failure,

extern "C" INTERFACE_API int XSPEC_rs232_getFirmwareAverage(HANDLE comport);

// return the trigger mode

// -1==failure,

extern "C" INTERFACE_API int XSPEC_rs232_getMode(HANDLE comport);

// 4 data structure stored in the flash. return one of them, use index (0,1,2,3) to indicate which one.

// pre-allocate the space for p_data, it is 128bytes. cast to the structure defined in myusb.h

// -1==failure, 0==success

extern "C" INTERFACE_API int XSPEC_rs232_get_SpecConfig0123(HANDLE comport,char *p_data, int index);

// 4 data structure stored in the flash. return one of them, use index (0,1,2,3) to indicate which one.

// create the data for p_data, it is 128bytes. use get_SpecConfig() to get the entire structure. Modify the one you need

// to change and put the data structure back using this function.

// return number of bytes sent, -1==failure,

extern "C" INTERFACE_API int XSPEC_rs232_write_SpecConfig(HANDLE comport,char *p_data, int index);

//only update the setting0 in memory, not update the flash

extern "C" INTERFACE_API int XSPEC_rs232_update_config_mem(HANDLE comport,char *p_data);

//retrieve the 4 data structures stored in the flash and save them locally.

// 0==success, -1==failure.

extern "C" INTERFACE_API int XSPEC_rs232_get_SpecConfig(HANDLE comport);

// dump the  4 data structures using printf

// // 0==success, else==failure

extern "C" INTERFACE_API int XSPEC_rs232_print_config(HANDLE comport);

// pre-allocate BUF_SIZE for the data_array, id is an unique ID to distingusish multiple users accessing the same device.

// passed integ_time help determine how long to wait for the data in usec; timeout time=integ_time/1000+1000ms

// pass data to read, pixels*2 + 2 for crc.

// return number of bytes received. <=0 indicates error.

extern "C" INTERFACE_API int XSPEC_rs232_measure(HANDLE comport,unsigned short *data_array, unsigned short bytes_to_read, unsigned int id, unsigned int integ_time);

// pre-allocate BUF_SIZE for the data_array, id is an unique ID to distingusish multiple users accessing the same device.

// passed integ_time help determine how long to wait for the data. in usec; integ_time/1000+1000ms

// pass rs_start and rs_end 36-3683 for normal operation. SAM3U

// size=(DATA_START+rs-end-rs_start+1)*2 +2

// return number of bytes received. <=0 indicates error.

extern "C" INTERFACE_API int XSPEC_rs232_measure_partial(HANDLE comport,unsigned short *data_array, unsigned short size, unsigned int id, unsigned int integ_time, unsigned short rs_start, unsigned short rs_end);

extern "C" INTERFACE_API int XSPEC_rs232_pwm_start(HANDLE comport, int which);

extern "C" INTERFACE_API int XSPEC_rs232_pwm_stop(HANDLE comport, int which);

// stop transmission,

extern "C" INTERFACE_API int XSPEC_rs232_abort(HANDLE comport);

// Restart the device. goto address "0"

extern "C" INTERFACE_API int XSPEC_rs232_reset(HANDLE comport);

extern "C" INTERFACE_API int XSPEC_rs232_measure_stop(HANDLE comport);

extern "C" INTERFACE_API int XSPEC_rs232_get_SpecProg(HANDLE comport, unsigned short page);

extern "C" INTERFACE_API int XSPEC_rs232_upgrade(HANDLE comport);

//only supported with DAC installed. adjust the DAC to reach average count of 90 for the dark signal

extern "C" INTERFACE_API int XSPEC_rs232_adjust_bias(HANDLE comport);

extern "C" INTERFACE_API int XSPEC_rs232_light_control(HANDLE comport,unsigned char lamp_num, unsigned char lamp_onoff);